## DBS
Deep Brain Stimulators
Vendor | Model | Field Strength | Interrogation | Conditions | Time

## VNS
Vagus Nerve Stim
Vendor | Model | Field Strength | Misc | Conditions | Time
Liva Nova VNS Therapy | Demipulse 103 AspireHC 105 AspireSR 106 Sentiva 1000 | 1.5 and 3 | Implanted Left Chest higher than 4th rib. | See Website for up to date specific parameters: https://vnstherapy.com/healthcare-professionals/mri
| All Pulse 102 Pulse Duo 102R DemiPulse Duo 104 Demipulse Duo 103 AspireHC 105 AspireSR 106 Sentiva 1000 | 1.5 and 3 | Implanted Left Chest lower than 4th rib. | See Website for up to date specific parameters: https://vnstherapy.com/healthcare-professionals/mri

## Sleep Apnea
Sleep Apnea
Vendor | Model | Field Strength | Interrogation | Conditions | Time
Inspire | 3028 | 1.5 |  | See Vendor specific Conditions for Head/Neck and Extremity scans only using Transmit Receive only coil.
| 3024 | NOT SAFE

## Bladder Bowel
Bladder Bowel
Vendor | Model | Field Strength | Information | Conditions | Time
Medtronic | 97810 Interstim Micro Sure Scan | 1.5 and 3 | With Lead 978A1 | Check Eligability with Remote Check Medtronic website for current parameters | Immediate
| 3058 Interstim II | 1.5 and 3 | With Lead 978B1 | Check Eligability with Remote Check Medtronic website for current parameters
| 3058 Interstim II | 1.5 Head T/R | Any Interstim Lead | Head Eligible Only
| 3023 | 1.5 Head T/R | Any Interstim Lead | Except with certain serial numbers
| 7427T Interstim Twin | NO MRI Scans
AXIOS | 1101 | 1.5T | With 1201 or 2201 leads | Device must pass MR readiness check. Stimulation Turned OFF. 2500 gauss/cm 200 T/m/s per axis Any receive only coil Normal Operating Mode Whole Body SAR equal to or less than 2W/kg Maximum 30 mins total scan with a 5 minute wait

## Spinal Cord Stims
Spinal Cord Stimulator
Vendor | Model | Field Strength | Information | Condition Status | Notes | Website Look-up Link
Boston Scientific | Preceision Montage SC-1200 | 1.5 | Leads will determine exact conditions | Whole body Conditional | Please see website for specific criteria and conditions | https://www.bostonscientific.com/elabeling/us/en/home/healthcare-professionals.html
| Precision Novi and Precision Plus SC-1140 |  |  | NOT MRI Conditional
| Precision Spectra SC-1132 | 1.5 | Leads will determine Conditional status | Head Only T/R Coil Only | Please see website for specific criteria and conditions
| Spectra WaveWriter SC-1160 | 1.5 | Leads will determine Contional status | Head Only T/R Coil Only | Please see website for specific criteria and conditions
| WaveWriter Alpha SC-1216, 1232,1416, 1432 | 1.5 | Leads will determine Contional status | Whole body Conditional | Please see website for specific criteria and conditions
Medtronic | 97702, 97712, 97713, 97714, 97715, 97716, 977005, 977006 | 1.5 | Remote control will determine condtionality. | Potential: Full-Body, or Head Only | Please see website for specific criteria and conditions | https://manuals.medtronic.com/manuals/main/en_US/home
| 37701, 37702, 37703, 37704, 37711, 37712, 37713, 37714, 7479, 7479B, 7427,7427V, 7425 | 1.5 | Remote control will determine condtionality. | Potential: Head Only | Please see website for specific criteria and conditions
Nevro | NIPG1000, NIPG1500, NIPG2000, NIPG2500 Senza | 1.5 | Leads will determine conditional status and exact scan parameters | Potential: Full-Body, or Head Only | Please see website for specific criteria and conditions | https://nevro.com/English/us/patients/patient-manuals/default.aspx
St Jude/Abbott | Proclaim DRG 3664 | 1.5 | With Lead MN10450-50/A 50cm only | Head and Extremity only with T/R coil only | Please see website for specific criteria and conditions | https://www.neuromodulation.abbott/us/en/hcp/provider-resources/mri-support.html
| Proclaim XR 5 3660, XR 7 3662 | 1.5 | With Octrode Lead 3186 60Cm only | Whole Body | Please see website for specific criteria and conditions
| Prodigy 3772 | 1.5 | With lead Octrode or Penta 60cm only | Head and Extremity only with T/R coil only | Please see website for specific criteria and conditions
| Protégé 3771 | 1.5 | With lead Octrode or Penta 60cm only | Head and Extremity only with T/R coil only | Please see website for specific criteria and conditions